Text S3 MALDI-TOF Settings.  The instrument and program settings for these experiments were as follow: General: Flex-Control Method- RP_pepmix.par. Processing: Flexanalysis Method- none, Biotools MS method- none. Laser Power: 30-38 % Sample Carrier: none.  Spectrometer: On, Ion Source 1- 19.0 0 mV, Ion Source 2- 16.40 mV, Lens- 9.45 mV, Reflector 20.00, Pulsed Ion Extraction- 190 ns, Polarity- Positive. Matrix Suppression: Deflection, Suppress up to: m/z 500.  Detection: Mass Range- 500-1600, Detector Gain- Reflector 3.7X.  Sample Rate- 2.00 GS/s, Mode- low range, Electronic Gain-Enhanced, 100 mV. Real time Smooth- Off. Spectrometer, Size: 81040, Delay 42968. Processing Method: Factory method RP_2465. Setup:  Mass Range- Low. Laser Frequency- 20 Hz, Autoteaching- off.  Instrument Specific Settings: Digitizer- Trigger Level- 2000 mV, Digital Off Linear- 127 cnt, Digital Off, Reflector- 127 cnt. Detector Gain Voltage Offset, Linear- 1300 V, Reflector- 1400 V. Laser Attenuator, Offset -12 %, Range- 30 %, Electronic Gain Button Definitions, Regular: 100 mv (offset lin) 100 mV (offset ref) 200mV/full scale. Enh: 51 mV (offset lin), 51 mV (offset ref) 100 mV/full scale. Highest: 25 mV (offsetlin) 25 mV (offset ref) 50 mV/full scale. Calibration:  Calibration was accomplished using angiotensin II as an external standard. Zoom Range +-1.0%, Peak Assignment Tolerance- User Defined-500 ppm.