| Temp. (oC) | DNA | Robustness attributes |
| --- | --- | --- |
| 122018 ASSAY |  |
| 57.5 | B. pseudomallei | Samples amplified ~1 CT later than at 60oC |
| 60 | B. pseudomallei | Best temperature; most consistent amplification |
| 62.5 | B. pseudomallei | Highest failure rate; amplified samples had precision similar to 60oC |
| 57.5 | B. thailandensis-like | Best temperature; most consistent amplification |
| 60 | B. thailandensis-like | Comparable success rate to 57.5oC; samples amplified ~1 CT later than at 57.5oC |
| 62.5 | B. thailandensis-like | High failure rate; much higher CTs than at 57.5 oC and 60oC |
| 266152 ASSAY |  |
| 57.5 | B. pseudomallei | Similar to 60oC; lower precision and samples amplified ~1 CT later than at 60oC |
| 60 | B. pseudomallei | Best temperature; most consistent amplification |
| 62.5 | B. pseudomallei | Similar to 60oC; replicates amplified ~1CT later than at 60oC |
| 57.5 | B. thailandensis-like | Highest failure rate; substantial loss of precision cf. 60oC |
| 60 | B. thailandensis-like | Best temperature; most consistent amplification |
| 62.5 | B. thailandensis-like | Higher failure rate cf. 60oC; much higher CTs than at 60oC |